Stereo Microscope

MZ41 SERIES

User Guide

Micro-shot Technology Limited

www.m-shot.com

Congratulations and Thank You for your purchase of the

Microscope MZ41

This precision instrument has been designed to require a minimum of optical and mechanical maintenance. Its excellent design assures years of high quality, reliable service. We recommend you read this entire manual carefully before beginning to use the instrument.

WARNING

PLEASE DISCONNECT PLUG FROM MAINS SOCKET BEFORE REMOVING THE BOTTOM PLATE, OPENING THE LAMP REPLACEMENT DOOR, OR REMOVING THE LAMP HOUSE

WARNING

DO NOT EXPOSE THE INSTRUMENT TO HIGH TEMPERATURES OR HUMIDITY.

AVOID USING THE INSTRUMENT IN EXTREMELY DUSTY LOCATIONS

OPERATING TEMPERATURE 5℃ TO 35℃

OPERATING HUMIDITY: 20% TO 80% （25℃）

CAUTION:	NEVER IMMERSE THE INSTRUMENT IN WATER OR SOLVENT

CAUTION:	DO NOT PUT ANY FOREIGN OBJECTS IN THE FRAME OR INTO ANY MOVING MECHANICAL PARTS

Table of Contents

Getting Started	3

System configuration	4

Technical Specifications	5

Component……………………………………………… 6

Setup Instructions	8

Basic Operation	9

Replacing reflecting Light Lamp	10

Replacing transmission Light Lamp…………………….. 11

Replacing the Fuse	12

Preventative Maintenance	12

Getting Started

This reference guide is written for college students and those who have experience on microscope operating. It contains basic operation steps without any rationale. Please read the instruction beforehand and set up the microscope according to the introduced order and method. Read the instruction once again before using to get familiar with it.

System configuration:

Technical specifications:

- Two objectives configurable, objective tube rotate to change power

- eyepiece tube inclination of 45°，360°rotatable, interpupillary distance adjustable (55mm~75mm)

- left eyepiece tube oxyoptex adjustable (-5~+5)

- reflected/transmission illumination, brightness adjustable

- Arm-type body or column body

- optical system technical configuration:

| Objective | 1X,3X | 2X,4X |
| --- | --- | --- |
| Eyepiece | WF10X | WF15X | WF10X | WF15X |
| Magnification | 10X,30X | 15X,45X | 20X,40X | 30X,60X |
| Field range(mm) | Φ20, Φ6.7 | Φ15,Φ5 | Φ10, Φ5 | Φ7.5, Φ3.8 |
| W.D(mm) | 88 |

Component:

Set-up instruction:

1. Remove all parts from their packing materials and retain the packaging in the event you need to transport the product.

2. Loosen the body locking screw and install the optical body to the frame after removing its dust cap. Tighten optical body with screw. 
DO NOT RELEASE THE OPTICAL BODY FORM YOUR GRASP UNTIL IT IS FIRMLY SECURED TO THE FRAME.

3. Remove the dust caps of the eyepiece tubes in the Optical Body and insert both eyepieces into the optical body.

4. Connect the power cord to the instrument and to a suitable power supply.

Basic Operation

1. Illumination Controls

As model built-in illumination, Connect the power cord to a suitable power supply, see figure 1. Turn on the power switch, and adjust the transmitted light brightness control. When turn on the reflected light brightness control, reflected light source is switched on, the angle of halogen lamp with the reflector cup can be adjusted.

2. Focusing Controls

Focusing adjustment is accomplished by using the focus knob located comfortably on each side of the frame. This arrangement allows for easy, precise adjustment without drift or discomfort. Turning either of the focus control knobs will raise or lower the optical body.

3. Interpupillary and oxyoptex Adjustments, see figure 3.

- 1) oxyoptex adjustment proper correction for individual vision is accomplished via the oxyoptex adjustment located at both eyepieces. Put the specimen on the stage, observe with your right eye; adjust the coarse/fine focusing hand-wheel until you can see the specimen clearly. Change to left eye, this time adjust the oxyoptex adjustment loop but not the focusing hand-wheel until a clear image appear.

2) Proper interpupillary distance, or the distance between eyepieces, is crucial to the comfort of the user. Adjusting the interpupillary distance is accomplished through a “turning” action of the optical head, at figure 3, allowing for quick and easy adjustment.

Replacing reflected (upper) Light

1. Disconnect the power plug and wait till the bulb is cool. Loosen the optical body set screw, and remove the optical body form the frame.

2. Loosen the lamp cover, the halogen bulb with the reflector cup can be removed, Insert the new one into the same fixture. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life.

3. Reinstall the lamp house. Switch on the power, adjustment brightness of illumination, check normal or not.

Replacing the Transmitted (lower) Light Lamp

1. Halogen lamp illumination, see figure 4
Disconnect the power plug and wait to be sure the bulb is cool. the halogen lamp socket is located in the base. Remove the slide holder and the table-board. Pull the old lamp straight out of the socket. Insert the new one into the lamp socket. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life. Switch on the power, adjustment brightness of illumination, check normal or not.

2. Fluorescent lamp illumination
Turn off and unplug the microscope. Loosen the screws, see figure 5. remove the old fluorescent lamp from the lamp socket. Insert the new one into the socket. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life. Reinstall the chassis. Switch on the power, check normal or not.

Replacing the Fuse

Turn off and unplug the microscope. The fuse socket is located on the back of base, see figure 5. You may unscrew the fuse house and exchange a new fuse, and then screw the fuse house again.

Preventative Maintenance

1. Cleaning frame and base

Disconnect the plug from mains socket before cleaning. Clean the frame and base with a soft cloth moistened with a mild detergent solution. Be sure the instrument is dry before using.

2. Cleaning optical parts

Microscope eyepieces and objectives are coated.

They should not be wiped while dry as dirt or dust may scratch the coating. It is best to remove parts from the frame prior to cleaning. Always blow loose dust away first. Use cotton swabs or lens tissue moistened with a lens cleaner or a small amount of alcohol, then wipe the surface clean with a good quality lens tissue. Solvents such as Xylene should NOT be used as cleaner.

# Text boxes

Base

Frame

Lens body

Eyepiece

1

2

3

4

11

5

6

7

8

9

10

1. eyepiece   2. oxyoptex adjustment loop  3. body lock screw  4. W halogen with reflected ring (with lampshade)  5. focusing hand-wheel  6. power switch  7. reflected light brightness knob  8. transmission light brightness knob  9. specimen holder  10.stage board  11. objective tube

Figure 1

1

Figure 2

- lock hand-wheel

1. oxyoptex adjustment ring 2. Objective tube

Figure 3

1

2

3

4

7

6

5

Figure 4

1. lamp cover 2. body lock screw 3. halogen with reflected ring 4. power switch  5. operation stage 6. Halogen 7. Specimen holder

1

2

3

4

5

Figure 5

1. screw   2. Fluorescence bulb 3 light socket 4 power switch 5 fuse socket